CANDY
EXOGAM
Efficiency calibrations and comparison
Ozge Aktas

---
Depends on the position
Low energy region
High energy region
Relative intrinsic efficiencies
conclusion
2
3
4
1
5

---
EXOGAMpositions
F2: Q=135°;F=180°
F3: Q=135°;F=45°
F4: Q=135°;F=90°
F5: Q=135°;F=270°
F6: Q=90°;F=0°
F7: Q=90°;F=315°
F8: Q=90°;F=270°
F9: Q=90°;F=225°
F10: Q=90°;F=180°
F11: Q=90°;F=135°
F12: Q=90°;F=90°
F13: Q=90°;F=45
14
Target
F12
F4
F3
F5
F13
F6
F7
F8
F11
F10
F2
F9

---
Efficiency of crystals , clover and all at the end of second experiment without Addback
%varepsilon _{p} = %varepsilon _{ %OMEGA } %varepsilon _{intrinsic}
Each crystal has same radius, and d should be equal for each of them which is target distance 14 cm
%varepsilon _{ %OMEGA } = {1} over {2%pi} (1- {d} over sqrt{ { d^{2}+ r^{2}}} )
The efficiency for each crystal is different although they supposed to have the same solid angle and similar intrinsic efficiency
There is a systematically different efficiencies for two different angle 90 and 135 degree.
Two possibility :
source is not in the correct position and solid angle is different for 90-135 degree
source is in good position but the 90degree clovers are not placed well.

---
Systematically 90degree clovers has low efficiency in 1408 keV source line (squares- 135 degree, circles 90degree

---
Intrinsict efficiency- relative
2
The idea is to compare the intrinsict efficiency of each crystal.
One way to be sure about the target distance 14 cm is to compare high energy 1.4 MeV peak in different angles.
But to make such comparison, we have to make assumptions that each crystal has a same intrinsic efficiency.
To find out relative intrinsic efficiency we can use 1460 keV line from background.
This g-ray is isotropic and coming to crystals in every angle. Therefore the efficiency of this line is not depend on the positioning of the clovers.
F12
F4
F3
F5
F13
F6
F7
F8
F11
F10
F2
F9

---
40P line 1460 keV is isotropic so for each crystal we should have a same count
Crystals doesn’t have same intrinsic efficiency
Crystal in 135 degree (squares) has better intr. Efficiency then 90degree ones
Around 15 % efficiency differences between two groups
Now compare this with 1408 keV line from source

---
1408 source
Effcor = (eff/N1460)*N1460max
1408/1460ratio
After 1408/1460 keV ratio correction
The effect of the degree dependence is less visible.
Its hard to say target position is wrong.
Different efficiency mostly depend on internal efficiency of crystal.

---
If all crystal has the same efficiency with crystal 2C (highest efficiency one)
With the Addback factor we reach 7% efficeincy in 1.4 MeV

---
Low energy
region
According to Emmanuel , there is a material between detectors and source. He saw an absorption in the low energy region.
- The highest efficiency should be around 121 keV
- While we can see it is at 244 keV.
- Threshold is generally less than 100 keV
- Some crystals are more damage than the others

---
Clover 3 @135 degree
Check the threshold vs absorption

---
For absorption ratio

---
Run425 after e805, Eu source
not useful there is no time alignment.
Even 3 clover have 0 ts all the time

---
F12
F4
F3
F5
F13
F6
F7
F8
F11
F10
F2
F9

---
Week 2
Second part (06 Nov.)
-60Co efficiency run comparison for the source position
-target/source position calculation
Effect of MUST2 trapozeid in low energy region

---
60co results run 568
Same trend as 152Eu source, the efficiency in 90 degree is lower than 135 degree.
Blue points for 1332 keV
Red point for 1172 keV
Square is 135deg
Dot is 90 deg

---
target distance (d) effect in solid angle
According to my understanding target can change position in Z direction
%OMEGA= 2%pi (1- {d} over { sqrt{ d^{2}+ r^{2}}} )
%varepsilon _{ %OMEGA } = {1} over {2%pi} (1- {d} over sqrt{ { d^{2}+ r^{2}}} )
EXOGAM :
Chaque cristal mesure 6 cm de diamètre à l’arrière (ils sont biseautés à l’avant pour assurer la compacité de l’ensemble et les faces ont un diamètre de 4.5 cm) et 9 cm de long. Le volume utile de détection est donc d’environ 250 cm3.
→
r=2.25 cm
d
r
Point source
Detector
Ω
Exogam 90
Exogam 135
Z
Beam direction

---
Change in the z direction calculation
Procedure :
Calculate the solid angle efficiency for both angles in different distance shifts
Two different method possible after this point:
  1. find two crystal in different angles that has the same counts for 1460 keV background peak. So the intrinsic efficiency of these two crystal is comparable
      1. group : clover7 crystal B (7B)in 90 degree vs clover 5 crystal B (5B)in 135.
      2. Second group: 8C in 90 vs 2B in 135
  2. Check the efficiency in 1408 keV or 60Co source
  3. Correct the 1408 kev efficiency according to solid angle
  4. Fit the corrected eff of two coupled crystals with pol1
  5. Using the slope of the pol1 find the zero crossing which will be target distance
  1. Check the efficiency in 1408 keV or 60Co source for all crystals
  2. Correct the 1408 kev efficiency according to solid angle
  3. Fit the corrected eff of crystals with pol1
  4. Using the slope of the pol1, find the zero crossing which will be target distance

---
Z / cm	90o	135 / +Z direction	135o / -Z direction
0.25	14.00	14.18	13.82
0.5	14.01	14.36	13.65
0.75	14.02	14.54	13.48
1	14.04	14.72	13.31
1.25	14.06	14.91	13.15
1.5	14.08	15.10	12.98
1.75	14.11	15.29	12.82
2	14.14	15.48	12.66
2.25	14.18	15.67	12.51
2.5	14.22	15.87	12.36
2.75	14.27	16.06	12.21
3	14.32	16.26	12.07
Exogam 90
Exogam 135
Z
Beam direction
14 cm
Target move to
 -z direction
Target move to
+z direction

---
Since we found that our 135 degree detector has better efficiency than the 90 degree detector,
 the source is probably moved in -Z direction so
source is closer to 135 degree detectors with respect to 90 degree ones
Now we will use the pairs that we choose to find the shift in the target position
Pairs: 7B 5B and 8C 2B
Eff / 90degre	90o	Eff / 135 +Z	135 / +Z direction	Eff / 135 -Z	135o / -Z direction	Z / cm
2.016E-03	14.00	1.967E-03	14.18	2.067E-03	13.82	0.25
2.014E-03	14.01	1.919E-03	14.36	2.119E-03	13.65	0.5
2.011E-03	14.02	1.872E-03	14.54	2.172E-03	13.48	0.75
2.006E-03	14.04	1.826E-03	14.72	2.226E-03	13.31	1
2.001E-03	14.06	1.782E-03	14.91	2.281E-03	13.15	1.25
1.994E-03	14.08	1.738E-03	15.10	2.338E-03	12.98	1.5
1.986E-03	14.11	1.696E-03	15.29	2.395E-03	12.82	1.75
1.977E-03	14.14	1.655E-03	15.48	2.454E-03	12.66	2
1.967E-03	14.18	1.615E-03	15.67	2.513E-03	12.51	2.25
1.955E-03	14.22	1.577E-03	15.87	2.574E-03	12.36	2.5
1.943E-03	14.27	1.539E-03	16.06	2.635E-03	12.21	2.75
1.930E-03	14.32	1.502E-03	16.26	2.697E-03	12.07	3

---
1460 COUNT by crystal
Pairs: 7B 5B
and
8C 2B
1408 efficiency
Clover 5 and 7 is better to compare, since there could be chamber effect in clover 8 which is directly under the target holder

---
slope from5B and 7B crystal comparison
Z distance
slope
5B vs 7B data for different distance
1408keV eff / solid angle efficiency
5B at
135 degree
7B
At 90 deg
1.49 cm

---
0.25 cm interval Z movement
crystals
1408keVeff/solidangle
slope of the trendlines
z distance shift in cm
slope from pol1 eqn
2.28 cm

---
Z 2.25 cm
1408keV eff / solid angle efficiency
The target is 2.28 cm shifted to the -Z direction
Z 1.5 cm
1408keV eff / solid angle efficiency
Exogam 90
Exogam 135
+Z
Beam direction

---
If the source has equal distance to each clover, the efficiency curves will be like this.
This is the results of efficiency curves if we move the target 2.28 cm to -Z direction.

---
All crystals after position correction

---
MUST2 absorption
121 keV efficiency
Square is at the end of second experiment
Dash is at the end of first experiment

---
Square with Must2
Star without Must2
For 135 degree detectors
244/778 keV efficiency
Not clear to say something about absorption effect.

---
Run425 after e805, Eu source without MUST2 detectors